# Analyzing exercise and heart rate:  ECG project

## Objective

The objective of this project is to develop and ECG recording and analysis system for  investigating  the effects of excersise on heart rate and the cardiac cycle.  The technical tasks are to:

- Develop LabView VI’s for data acquisition and analysis

- Develop a signal conditioning circuit for ECG pre-processing

- Develop resonable measures for heart rate analysis from the ECG signal

The scientific tasks are to design and complete experiments to address the following questions:

- Does resting heart rate vary with constant factors, such as age, height or gender?

- Does heart rate vary with exercise? What is the relation?

- Demonstrate another factor which varies heart rate?

- Is Einthoven’s triangle verifiable with your system?

## Materials

- Surface ECG electrodes

- LabView and LabPC+ DAQ board

- Electronic components & circuit board

- Electronic test equipment

## Procedures

Each team is to develop a full-fledged, computer-based, biomedical instrumentation system to record and analyze the ECG.   The system will be used to investigate the  scientific questions listed above involving the assessment of heart rate for different subjects under different conditions.   Each group will design experiments and collect the appropiate data (using class members as subjects) to reach some type of data-based conclusion for each question.

The responsibilities of each team member are described below. While you are responsible for completing your individual tasks, you are also responsible for collaborating with other team members on their tasks. Remember the lab report is a combined grade for all members of the group.

### Hardware Engineers

- Design and prototype a filter/amplifier with an overall gain that is 1000.  Include active high and lowpass filtering (i.e a bandpass filter) in your design.  Select your high and lowpass cutoffs based on the expected freqeuncy content of the ECG (e.g., see Webster) and the expected frequency content of the your expected noise sources.  In particular, your design must reject motion artifact and electrode offset potentials (see Webster again)

- Test the final circuits and document with schematics and test results (use known input signals and present obtained outputs).  In addition to documenting the differential frequency response of your system, also measure the common mode gain for your circuit (note: this should be measured at a relevant freqeuncy!).

- Work with systems analyst on running and analyzing the experiments

### Software Engineers

- Design a LabView VI to sample the ECG data as required. Nominally, work with the hardware engineer to determine an appropriate sampling frequency.

- Decide on file structures for any input and output data files

- Develop a set of test data for verifying the algorithms and the programs

- Design and implement the data analysis functions required for calculating either instantaneous or averaged heart rates

- Test your program thoroughly

- Document your programs

- Work with systems analyst on running and analyzing the experiments

### Systems Analysts

- Develop and update the system block diagram

- Work with the team to describe the signals and the transformations at each stage of the system

- Design the experimental protocol required to answer the given questions

- Document the experiments and the data collection

- Design system-level, end-to-end tests of the entire system

- Take the lead role in data analysis

Reference Material:

Section 5.5 and Chapter 6 of Webster (1992), Medical Instrumentation: Application and Design.