# Estimating Muscle Force from the Surface EMG

## Objective

The objective of this project is to investigate the relationship between the force generated by a muscle and the muscle’s surface electromyogram (or EMG).  The technical tasks are to:

- Develop LabView VI’s for data acquisition and analysis

- Develop a signal conditioning circuit for EMG pre-processing

- Develop a reasonable measure that relates the EMG with muscle force.

The scientific tasks are to design and complete experiments to address the following questions:

- How does the EMG vary with muscle force?

- Does EMG vary with fatigue?

- Is the EMG a consistent and reliable estimate of muscle force?

- Is “trained” muscle more efficient than an untrained muscle?

## Materials

- Surface EMG electrodes

- Dumbbells

- LabView and LabPC+ DAQ board

- Electronic components & bread board

- Electronic test equipment

## Procedures

Each team is to develop a full-fledged, computer-based, biomedical instrumentation system to record and analyze the EMG following the broad guidelines given above.  The system will be used to investigate the scientific questions listed above involving the estimation of muscle force from EMG.  Each group will design experiments and collect the appropriate data (using class members as subjects) to reach some type of data-based conclusion for each question.

The responsibilities of each team member are described below. While you are responsible for completing your individual tasks, you are also responsible for collaborating with other team members on their tasks.

### Hardware Engineers

- Design and prototype a single-ended amplifier with the gain you decided upon in the Pre-Lab. Remember, you already have a preamplifier designed from the Intro Lab. Keep this in mind when designing your overall gain.  You may use / modify your original circuit if you feel it would be appropriate.

- Design and prototype a separate active band-pass filter with a cutoff frequencies from the Pre-Lab.

- Test the final circuits and document with schematics and test results (use known input signals and present the results obtained).

- Work with systems analyst on running and analyzing the experiments.

### Software Engineers

- Design a LabView VI to sample the EMG data as required. This requires sampling at a bare minimum of twice your high cutoff of the band-pass filter.

- Decide on file structures for any input and output data files

- Develop a set of test data for verifying the algorithms and the programs

- Design and implement the data analysis functions required for calculating either the moving average or the root-mean-square muscle activation (see equations below).

- Test your program thoroughly

- Document your programs

- Work with systems analyst on running and analyzing the experiments.

### Systems Analysts

- Develop and update the system block diagram

- Work with the team to describe the signals and the transformations at each stage of the system.

- Design the experimental protocol required to answer the given questions

- Design system-level, end-to-end tests of the entire system

- Take the lead role in data analysis.

EMG equations